## Slide 1
Geranyl acetate C12H20O2

## Slide 2: Mass Spectral Libraries An Ever-Expanding Resource for Chemical Identification
Steve Stein
Mass Spectrometry Data Center
National Institute of Standards and Technology
Gaithersburg, Maryland, USA

## Slide 3
NIH/EPA Collection of Collections
Fales, Heller
Red Books 9-track Tape 300 Baud Modem
To NIST
PC-XT Version
Evaluated Library
AMDIS
Exact m/z
Peptides
To EPA
Cincinnati
Budde
Structures
Begin Manual Evaluation
Tandem MS
GC Retention
2000’s
1990’s
1980’s
1970’s
Evolution of the NIST MS Library
2010’s

## Slide 4: NIST/EPA/NIH MS Library
Chart: Libraries Distributed/Year
| '88 | '89 | '90 | '91 | '92 | '93 | '94 | '95 | '96 | '97 | '98 | '99 | '00 | '01 | '02 | '03 | '04 | '05 | '06 | 7 | '08 | 09 | '10 | '11
 | 1288 | 1527 | 1679 | 1980 | 2206 | 2238 | 3699 | 3546 | 3190 | 3074 | 3090 | 3462 | 3139 | 2795 | 3177 | 3458 | 3832 | 3955 | 4446 | 4547 | 4517 | 4689 | 4156 | 4159

## Slide 5: Library Growth

## Slide 6: Mass Spectra
Flux of Ions at Discrete Masses
  Actually Mass/Charge (m/z)
  Mass reflects elemental composition
    Methane = CH4 = 16 nominal = 16.0303 u
    Isotope = 13CH4 = 17.0348 u (1%)
  Perfect spectrum – intensity vs chemical formula
Ion is a ‘charged’ molecule (interconnected atoms)
  # protons ≠ # electrons
  Ions controllable with electric/magnetic fields

## Slide 7: Mass Spectrometers
GC/MS Electron Ionization
LC/MS: Electrospray/MALDI Ion Trap/Collision Cell/..

## Slide 8
Samples are often very complex
GC/MS of Environmental Sample
Retention Time
Signal
Intensity
Notes: I expect that you do not plan to add reference spectra from such chromagrams.
The spectrum of a target compound can be influenced by the presence other others eluting nearby. Complex matrices present special problems. Not only can actual peaks in a spectrum be obscured and spurious peaks added, but the spectrum extraction process itself can seriously distort spectra.

## Slide 9: LC-MS/MS of Protein Digest

## Slide 10: Goal: Chemical Identity
Elemental Composition (measurable by MS)
Chemical Structure (invisible to MS)
Notes: I promised I wouldn’t talk much about chemistry – but I must talk about chemicals – as mathematical objects

## Slide 11: Reveal Structure as Spectrum: A Mass “Fragmentogram”

## Slide 12: Structure/Spectrum Space examples of structures with similar spectra

## Slide 13: Mass Spectra Reproducible Over Time
O’Neal et al. Anal. Chem. 1951
NIST 2012
A mass spectrum is a property of an ion
Notes: .. Why do I say that <next slide>

## Slide 14: Molecular Fingerprints
VX
HD
GB
Notes: Since depend only of molecule, are reproducible – can serve as fingerprints

## Slide 15: Spectra Can be Interpreted, Not Predicted
MS Interpreter
?
Notes: So – need libraries

## Slide 16: Library Search
“Fingerprint” Identification
  Identify compound by matching spectrum to library spectrum
Compare Query to Library Spectra
  Derive ‘score’ for each library spectrum
    Reflect likelihood both spectra are from same compound
  Arrange results by score
    Create ‘Hitlist’
Notes: 2nd bullet – more specifically

## Slide 17: Traditional Library Search
Query
Spectrum
Library
Spectrum
Hit
List
Score
Histogram
2011 Version - 213K EI, 5K CID, 71K RI Compounds
Search List

## Slide 18: Non-Traditional Peptide Spectrum Search
Query
Spectrum
Library
(Consensus)
Spectrum
Hit
List
Score
Histogram
For Protein Inference and Quality Monitoring
Spectrum List

## Slide 19: Spectrum Similarity Score
Each spectrum expressed as a vector
Cosine of Angle between Query and Library Spectra
Q, L = weight (abundance, mass) : for each peak
weight (abundance, mass)
  Abundance
  Abundance * m/z
  Certainty * Abundance
  …
S.E. Stein, D.R. Scott "Optimization and Testing of Mass Spectral Search Algorithms for Compound Identification”, J. Amer. Soc. Mass Spectrom., 5, 859-866 (1994)
Notes: The basic form of the ‘correlation function’ – it can be viewed as a normalized ‘dot product’ or related to the cosine between the ‘angle’ between spectra when expressed as vectors. Optimal peak weighting can improve results somewhat at placing more emphasis on less common peaks (higher masses, for example, are more discriminating that lower masses for identification).

## Slide 20
Peaks are Highly Correlated by mass and abundance
Differences in m/z
Relative Frequency
Big Peaks
Small Peaks
Medium Peaks
S.E. Stein and D.N Heller, J. Amer. Soc. Mass Spectrom. 2006, 17, 823-835.

## Slide 21: Score Confidence Level
How to Express Identification Certainty?
  Related to broad range of Identity problems
  Can it be quantitative?
Follow Bayes
  Follow changes in confidence
Bayesian Notation
  P ( ID is correct | Threshold Score )
Notes: Score – what we get; confidence – what we want
Even if experts are convinced – need way to convince others
Reverend Thomas Bayes – 1702 - 1761

## Slide 22: Bayes Rule*
Reproducible
Spectrum
Prior Probability:
Before Experiment
False Positive
Potential
Analyte is Identified Correctly
Final
Confidence
Starting Confidence
Change in Confidence
* Odds Version
P ( Score | ID)
P ( ID )
P ( FP )
P ( Score | FP)
X
Influence of Library Search
P ( ID | Score )
P ( FP | Score )
Notes: -End – OK – now that we know how to express identification confidence, what do we want to identify

## Slide 23: I. Prior Probability
Notes: A more famous US gov’t library

## Slide 24: How plausible is the ID?
Seen before under similar conditions?
Expert knowledge
  Expected, plausible, unlikely, impossible
Citations
  Google, ChemSpider, PubChem, MS Library, …
  Human Metabolite DB, Merck Index, ..
Weak link in Identification
  Most compounds in a library cannot be in sample
Notes: Sarin story, fulvene, plasma stories – need expert to find citations

## Slide 25: II. Spectrum Variability P( Score | ID )
False Negative Potential
  FN when Correct ID has low score
Spectra are Kinetic Properties of Ions
  But can vary due to instrument bias
  Include known variations in library
But, Spectra Can Vary For Other Reasons
  Low S/N
  Contaminants
  Instrument problem
  Chemical reaction before/after Ionization
Notes: Probability that a correct ID will yield high score – assume for now in library

## Slide 26
Instrument ‘Noise Signature’ 250 Hexachlorobenzene Spectra same instrument, calibration mix
Bars show quartiles
Notes: This shows instrument dependent variations for individual peaks. In this instrument, variability of abundance increased with increasing m/z ‘distance’ from the base peak. Variations of peaks at nearby m/z can be adequately described as a fixed fraction of the actual abundance. Other noise behavior is observed on other instruments or even under different analytical conditions.

## Slide 27: Typical Interlab Spectrum Variation
Notes: Raphael Mechoulam

## Slide 28: Energy Dependence
[M+H+K]2+
30
35
40
CollisionEnergy Setting

## Slide 29: Ion Source Decomposition
150 C
280 C

## Slide 30: III. False Positive Potential P( Score | FP )
Wrong Compound ID – High Score
Analyte Structure Is Unique
  But is it’s spectrum?
  NO – Cannot be shown beyond ‘reasonable doubt’
    Spectrum insufficient to ‘reconstruct’ molecule
  But, may be unique for all plausible compounds
    FP should be restricted to plausible compounds
Sometimes Only Class ID is Possible
  Even if retention time is matched
Notes: Most significant worry when analysis is confident may be present
Prior probability – hire expert, variability – do much work, fp potential – too difficult
If all imaginable compounds were equally likely to be present – we would be in trouble

## Slide 31
Science, Aug, 22, 2008
Asara et al.
Response to Comment on “Protein Sequences from Mastodon and Tyrannosaurus rex Revealed by Mass Spectrometry”
Large Libraries Can Show Uniqueness

## Slide 32: Varieties of FPs
Accidental
  Rare for good quality, informative spectra
Low information content
  Few reliable peaks
MS Class Identification
  Different compounds yield same set of ions due to structure similarity
Notes: Prior probability – hire expert, variability – do much work, fp potential – too difficult

## Slide 33: MS Specific Class ID

## Slide 34: MS Specific Class ID

## Slide 35: Same low mass ions (benzoyl)

## Slide 36: Random Match

## Slide 37: False Positives above 800

## Slide 38: Hit List Contains Structural Information
S.E. Stein "Chemical Substructure Identification by Mass Spectral Library Searching”, J. Amer. Soc. Mass Spectrom. 6 (1995), 644-655.

## Slide 39: Is Analyte In Library?
Ideal library contains all plausible compounds
P(correct) = P(in library) x P(correct, assuming in library)
      Estimating Probabilities of Correct Identification from Results of Mass-Spectra Library Searches, JASMS 5 p. 316 (1994)
P(in library) = f(score)
P(correct compound) = f(score – score next best ID)
Optional ‘Prior Probability’ correction compound is in other collections

## Slide 40: How to Handle Unknown Components?
Major Category: Plant “Metabolites”
Vetiver Oil

## Slide 41: Vetiver Oil Many components not identified by GC/MS
NIST Library Identified Manually Not Identified

## Slide 42: Unidentified Recurring Spectrum Library
From 3,700 pediatric urine samples. Derived 203 recurring spectra with same retention.
Working on blood, essential oils , biologic drugs
99% - RI 1200
69% - RI 1860
57% - RI 2504

## Slide 43: Rumsfeld Quadrants
Expected by Analyst
Unexpected
by Analyst
Identified
by Library
Not Identified
by Library
Recurrent Spectral Libraries
Concentration too low, not in library, …
Comprehensive Library
Known Knowns
Expected and found
Unknown Knowns
Not expected but found
Known Unknowns
Expected but not found
Unknown Unknowns
Not expected and not found
Target Library
Notes: Donald Rumsfeld inspired – feb 2002
There are known knowns. These are things we know that we know. There are known unknowns. That is to say, there are things that we know we don't know. But there are also unknown unknowns. There are things we don't know we don't know. Read more:http://www.brainyquote.com/quotes/quotes/d/donaldrums148142.html#ixzz1k2l6HO2k

## Slide 44: Our Pipeline